Protocol of Acetylcholinesterase Enzyme-based Assay for High-throughput Screening

| DOCUMENT: |  | AChE_ tox21_Version3.0 |
| --- | --- | --- |
| TITLE: |  | Protocol of Acetylcholinesterase Enzyme-based Assay for High-throughput Screening |

ASSAY RFERENCES:

| Assay Target | Assay Readout | Assay Provider | Toxicity Pathway |
| --- | --- | --- | --- |
| Acetylcholinesterase | Absorbance | NCATS | neuron |

QUALITY CONTROL PRECAUTIONS:

1. -The concentration of the substrate (acetycholine) should be less than the Km.

2. -The acetylcholinesterase used should also be optimized.

3. –The enzyme is plated into 1536-well plate first.

4. -Handle 1536 well plate black clear bottom plates carefully by sides

MATERIALS and INSTRUMENTS:

| Supplies/Medium/Reagent | Manufacturer | Vender/Catalog Number |
| --- | --- | --- |
| -BiorapTR | -Beckman Coulter | - |
| -Envision | -Perkin Elmer | - |
| -Colorimetric reagents | -AAT Bioquest | - 11400 |
| - Purified recombinant human AChE | - Sigma-Aldrich | - C168 |

PROCEDURE:

1 Assay Protocol

1.1 - Plate the enzyme in black-clear bottom 1536 well plate at 4uL for antagonist mode using a single tip plate dispenser (BiorapTR)

1.2 - Transfer 23nL of compounds from the library collection and positive control to the assay plates through pintool

1.3 -Incubate for 1 hr at room temperature

1.4 -Add 4 uL of cocktail solution (DTNB, acetylthiocholine) dye using a single tip plate dispenser (BiorapTR)

1.5 -Incubate at room temperature for 10 min in dark

1.6 -Read the absorbance intensity through Envision plate reader using protocol optimized