Protocol of ADRB2 Cell-based Assay for High-throughput Screening

| DOCUMENT: |  | ADRB2_TOX21_SLP_Version1.0 |
| --- | --- | --- |
| TITLE: |  | Protocol of CHO-ADRB2 Cell-based Assay for High-throughput Screening |

ASSAY RFERENCES:

| Assay Target | Cell Lines | Species | Tissue of Origin | Assay Readout | Assay Provider | Toxicity Pathway |
| --- | --- | --- | --- | --- | --- | --- |
| ADRB2 | CHO-K1 | Hamster | Chinese hamster ovary | Homogeneous Time Resolved Fluorescence (HTRF) | Cisbio (ThermoFisher Scientific) | Gs-coupled |

QUALITY CONTROL PRECAUTIONS:

1. Maintain cells below 85-90% confluence.

2. Feed or passage cells twice a week.

MATERIALS and INSTRUMENTS:

| Supplies/Medium/Reagent | Manufacturer | Vender/Catalog Number |
| --- | --- | --- |
| DMEM/F12 | ATCC | ATCC (30-2006) |
| Fetal Bovine Serum, Qualified | GIBCO (ThermoFisher Scientific) | GIBCO (26140) |
| Penicillin/Streptomycin | GIBCO (ThermoFisher Scientific) | GIBCO (15140) |
| Puromycin | GIBCO (ThermoFisher Scientific) | GIBCO (A11138-03) |
| Recovery Cell culture Freezing Medium | GIBCO (ThermoFisher Scientific) | GIBCO (12648) |
| 0.05% Trypsin-EDTA | GIBCO (ThermoFisher Scientific) | GIBCO (25300) |
| 1536-well white solid plates | Greiner Bio-One | Greiner Bio-One (789173-F) |
| MULTIDROP COMBI | ThermoFisher Scientific | ThermoFisher Scientific |
| BioRAPTR FRD | Beckman Coulter | Beckman Coulter |
| Envision Plate Reader | Perkin Elmer | Perkin Elmer |
| CAMP - GS DYNAMIC KIT | Cisbio (Perkin Elmer) | Cisbio (62AM4PEJ (100,000 tests)) |
| Isoproterenol hydrochloride (Agonist control compound) | Millipore Sigma | Millipore Sigma (I6504) |
| Propranolol hydrochloride (Antagonist control compound) | Millipore Sigma | Millipore Sigma (P0884) |
| 3-Isobutyl-1-methylxanthine (IBMX) | Millipore Sigma | Millipore Sigma (I7018) |
| Ro-20-1724 | Millipore Sigma | Millipore Sigma (557502) |

PROCEDURE:

1. Cell handling:

1.1. Media Required:

| Component | Growth Medium | Assay Medium | Thaw Medium | Freezing Medium |
| --- | --- | --- | --- | --- |
| DMEM/F12 | 90% | 90% | 90% | - |
| Fetal Bovine Serum, Qualified | 10% | 10% | 10% | - |
| Penicillin-Streptomycin (10,000 U/mL) | 1% | 1% | 1% | - |
| Puromycin | 10 µg/mL | - | - | - |
| Recovery Cell Culture Freezing Medium | - | - | - | 100% |

1.2. Thawing method:

1.2.1 Thaw a vial of cells in 9 mL of pre-warmed medium and seed them in T-175 flask at 2 x 106 cells.

1.3. Propagation method:

1.3.1 Trypsinize cells from the flask and centrifuge.

1.3.2 Add culture medium to the pellet and passage at 2-3 x 106 cells per T-225 flask.

2. Assay Protocol:

2.1. Harvest cells from the culture and re-suspend cells in assay medium at 0.25 x 106 cells/mL.

2.2. Dispense 1,000 cells/4 µL/well in 1536-well tissue treated white/solid bottom plates using Multidrop COMBI.

2.3. Incubate the assay plates for an overnight (18-20 hr) at 37°C and 5% CO2.

2.4. Transfer 23 nL of the test compounds from the libraries and positive control to the assay plates through Pintool.

2.5. Add 1 µL of stimulation mixture (Prepare a mixture of 100 µM IBMX and 25 µM Ro-20-1724 in stimulation buffer provided in the Cisbio kit) to all the wells of the assay plates using BioRAPTR FRD.

2.6. Centrifuge the assay plates at 1000 rpm for 15 sec.

2.7. Incubate the assay plates for 0.5 hr at room temperature.

2.8. Add 2.5 µL of cAMP-d2 working solution to all the wells of the assay plates using BioRAPTR FRD. (Prepare 20X stock solution of cAMP-d2 by reconstituting cAMP-d2 with 5 mL distilled water and then dilute one volume of 20X stock solution in 19 volumes of Lysis & Detection buffer.)

2.9. Add 2.5 µL of Anti-cAMP-Cryptate to all the wells of the assay plates using BioRAPTR FRD. (Prepare 20X stock solution of Anti-cAMP-Cryptate reconstituting Anti-cAMP-Cryptate with 5 mL distilled water and then dilute one volume of 20X stock solution in 19 volumes of Lysis & Detection buffer.)

2.10. Incubate the assay plates at room temperature for 1 hr.

2.11. Read the fluorescence intensity using Envision plate reader with an excitation wavelength at 340 nm and emission at two different wavelengths; 665nm and 615nm. Final data are expressed as ratio of 665nm/620 nm.

3. Assay Performance:

| ADRB2 (Agonist mode) | Online Screening (Mean ± SD) |
| --- | --- |
| AC50 | 3.32 ± 1.9 nM (n = 404) |
| S/B | 1.92 ± 0.13 |
| CV (%) ⃰ | 7.51 ± 0.98 (n = 406) |
| Z’ | 0.42 ± 0.13 |

⃰ CV values shown represent average of DMSO-only column from all the assay plates.