Protocol of ER-alpha-BG1 BG1 Cell-based Assay for High-throughput Screening

| DOCUMENT: |  | ER-alpha-BG1_TOX21_SLP_Version1.0 |
| --- | --- | --- |
| TITLE: |  | Protocol of ER-alpha-BG1 BG1 Cell-based Assay for High-throughput Screening |

ASSAY RFERENCES:

| Assay Target | Cell Lines | Species | Tissue of Origin | Assay Readout | Assay Provider | Toxicity Pathway |
| --- | --- | --- | --- | --- | --- | --- |
| Estrogen receptor alpha: full (Endogenous) | vMCF-7 (Misidentified as BG1) | Human | Breast carcinoma | Luciferase reporter | UC Davis | NR signaling |

QUALITY CONTROL PRECAUTIONS:

1. -Maintain cells below 85-90% confluence in culture medium

2. -For assay purpose, the cells should be cultured in assay medium with 10% charcoal stripped FBS for 5 days with alternate day medium changed to fresh medium

3. -Especially while in assay culture, the cells should not reach more than 85% confluence as they would become harder to detach if they reach over confluence

MATERIALS and INSTRUMENTS:

| Supplies/Medium/Reagent | Manufacturer | Vender/Catalog Number |
| --- | --- | --- |
| -MEM α medium | -Invitrogen | -Invitrogen / 12561 |
| -10% Premium Fetal Bovine Serum | -Atlanta Biologicals | -Atlanta Biologicals / S11150 |
| -Penicillin/Streptomycin | -Invitrogen | -Invitrogen / 15140 |
| -400mg/l G418 (Geneticin) | -Invitrogen | -Invitrogen / 10131 |
| -DMEM phenol red free - low glucose medium | -Invitrogen | -Sigma Aldrich / D5921 |
| -Charcoal stripped Fetal Bovine Serum | -Invitrogen | -Invitrogen / 12676 |
| -L-Glutamine | -Invitrogen | -Invitrogen / 25030 |
| -0.25% Trypsin-EDTA | -Invitrogen | -Invitrogen / 25200 |
| -1536-well white solid plates | -Corning | -Corning / 7464 |
| -MULTIDROP COMBI | -Thermo Electron Corporation | -Thermo Electron Corporation |
| -BioRAPTR FRD | -Beckman Coulter | -Beckman Coulter |
| -ViewLux Plate Reader | -Perkin Elmer | -Perkin Elmer |
| -ONE-Glo Luciferase Assay system | -Promega | -Promega / E6130 |
| -CellTiter-Fluor(TM) Cell Viability Assay | -Promega | -Promega / G6082 |
| -Recovery Cell culture Freezing Medium | -Invitrogen | -Invitrogen / 12648 |
| -Beta-Estradiol (Agonist control compound) | -Sigma Aldrich | -Sigma Aldrich / E8875 |
| -ICI 182,780 (Antagonist control compound) | -Sigma Aldrich | -Sigma Aldrich / I4409 |
| -Tetraoctyl ammonium bromide (Viability positive control compound) | -Sigma Aldrich | -Sigma Aldrich / 294136 |

PROCEDURE:

1. Cell handling:

1.1. Media Required:

| Component | Growth Medium | Assay Medium | Thaw Medium | Freezing Medium |
| --- | --- | --- | --- | --- |
| -MEM α medium | -90% | - | -90% | - |
| -DMEM phenol red free - low glucose medium | - | -90% | - | - |
| -Premium Fetal Bovine Serum | -10% | - | -10% | - |
| -Charcoal/dextran treated Fetal Bovine Serum | - | -10% | - | - |
| -Penicillin/Streptomycin | -100U/ml & 100ug/ml | -100U/ml & 100ug/ml | -100U/ml & 100ug/ml | - |
| -L-Glutamine | - | -2mM | - | - |
| -G418 (Geneticin) | -400mg/l | - | - | - |
| -Recovery Cell culture Freezing Medium | - | - | - | -100% |

1.2. Thawing method

1.2.1 -Thaw a frozen vial of cells in 9ml of pre-warmed medium and seed them in T175 flask at 2 million cells

1.3. Propagation method

1.3.1 -Trypsinize cells from the flask and centrifuge

1.3.2 -Add culture medium to the pellet and passage at 3-4 million per T-225 flask

2. Assay Protocol

2.1 -Harvest cells from the 5-day culture in assay medium and resuspend cells in assay medium.

2.2 -Dispense 4000 cells in 4uL/well into 1536-well tissue treated white/solid bottom plates using an 8 tip dispenser (Multidrop).

2.3 -Incubate the assay plates for 22hrs at 37C and 5% CO2.

2.4 -First 1uL of 10.0nM (final concentration) ICI-182,780 (ER-Antagonist) or assay buffer was added using two separate tips of a dispenser (BioRAPTR).

2.5 –Then transfer 23nL of compounds from the library collection and positive control to the assay plates by using a Pintool station.

2.6 -Incubate the assay plates for 22hrs at 37C and 5% CO2.

2.7 -After 21hrs of incubation, 1ul of CellTiter-Fluor(TM) Cell Viability Assay reagent was added using a single tip of a dispenser (BioRAPTR).

2.8 -Incubate the assay plates at 37C and 5% CO2 for 1hr.

2.9 -Measure fluorescence signal by ViewLux plate reader (Exposure time = 1sec).

2.10 -Then add 4ul of ONE-Glo(TM) Luciferase reagent using a single tip of a dispenser (BioRAPTR).

2.11 -Incubate the plates at room temperature for 30min.

2.12 -Measure luminescence signal by ViewLux plate reader (Exposure time = 30sec).

3. Assay Performance

| ERα-BG1 ICI-182,780 added | Online Validation Agonist (Mean ± SD) | Online Validation Viability (Mean ± SD) |
| --- | --- | --- |
| EC50 ICI-182,780-free β-estradiol | 0.019nM ±0.005 (n = 27) | NA |
| EC50 β-estradiol with 10.0nM ICI-182,780 | 0.94nM ±0.11 (n = 27) | NA |
| S/B | 8.79 ± 1.58 | 7.07 ± 0.20 |
| CV (%) | 15.72 * ± 3.53 (n = 20) | 6.39 * ± 0.51 (n = 21) |
| Z’ | 0.82 ± 0.03 | 0.85 ± 0.03 |

* CV values shown represent average of all assay plates excluding the top two concentration plates