Protocol of PGC/ERR Hek293T Cell-based Assay for High-throughput Screening

| DOCUMENT: |  | PGC/ERR_TOX21_SLP_Version1.0 |
| --- | --- | --- |
| TITLE: |  | Protocol of PGC/ERR Hek293T Cell-based Assay for High-throughput Screening |

ASSAY RFERENCES:

| Assay Target | Cell Lines | Species | Tissue of Origin | Assay Readout | Assay Provider | Toxicity Pathway |
| --- | --- | --- | --- | --- | --- | --- |
| PGC/ERR regulated genes | PGC/ERR Hek293T | Human | Embryonic kidney | Luminescence | NTP | Energy homeostasis |

QUALITY CONTROL PRECAUTIONS:

1. The cells should not be grown more than 80-85% confluence

2. Do not leave cells in Trypsin for more than 5 min at RT

MATERIALS and INSTRUMENTS:

| Supplies/Medium/Reagent | Manufacturer | Vender/Catalog Number |
| --- | --- | --- |
| DMEM, high glucose | Invitrogen | 11965-092 |
| HyClone® FBS | Hyclone | SH30071-03 |
| L-glutamine,200mM | Invitrogen | 25030-081 |
| Sodium pyruvate solution, 100 mM, | Invitrogen | 11360-070 |
| Penicillin/Streptomycin (antibiotic) | Invitrogen | 15140 |
| Puromycin | Invitrogen | A11138-03 |
| Recovery Cell Culture Freezing Medium | Invitrogen | 12648 |
| ONE-Glo Luciferase Assay | Promega | E6051 |
| CellTiter Fluor Cell Viability Assay | Promega | G6080 |
| ViewLux Plate Reader | Perkin Elmer | - |
| Multidrop | Thermofisher | - |
| BioRAPTR | Beckman Coulter | - |

PROCEDURE:

1. Cell handling:

1.1. Media Required:

| Component | Growth Medium | Assay Medium | Thaw Medium | Freezing Medium |
| --- | --- | --- | --- | --- |
| DMEM, high glucose | 90% | 90% | 90% | - |
| HyClone® FBS | 10% | 10% | 10% | - |
| Lglutamine,200mM | 4mM | 4mM | 4mM | - |
| Sodium pyruvate solution, 100 mM | 1mM | 1mM | 1mM | - |
| Penicillin/Streptomycin (antibiotic) | 1% | 1% | 1% | - |
| Puromycin | 1ug/mL | 1ug/mL | - | - |
| Recovery Cell Culture Freezing Medium | - | - | - | 100% |

1.2. Thawing method

1.2.1 Place 14 mL of pre-warmed thaw medium into a 15 mL conical tube

1.2.2 Remove the vial of cells to be thawed from liquid nitrogen and thaw rapidly by placing at 37°C in a water bath

1.2.3 Mix the entire content of the vial to 14 ml of pre-warmed medium and centrifuge to remove DMSO

1.2.4 Discard the supernatant and reconstitute the pellet using pre-warmed thaw medium

1.2.5 Transfer the cells to T75 flask

1.3. Propagation method

1.3.1 Detach the cells from the flask using Trypsin-EDTA (0.05%)

1.3.2 The cells in growth medium are re-seeded in T225 flask

2. Assay Protocol

2.1 Harvest cells and re-suspend in assay medium and adjust the required cell density

2.2 Dispense 2500 cells/5 µL/well into 1536 well tissue treated white plates using a Multidrop dispenser

2.3 Incubate the plates at a 37ºC, 5% CO2 incubator for 6 h

2.4 Transfer 23nL of compounds and positive control to the assay plate by a pin tool.

2.5 Incubate the assay plates at 37°C, 5% CO2 for 17.5 h

2.6 Add 1 µL of CellTiter-Fluor regent using a BioRAPTR

2.7 Incubate at 37ºC, 5% CO2 for 17.5 h

2.8 Read the fluorescence intensity in ViewLux plate reader

2.6 Add 4uL of OneGlo and incubate the plates at room temperature for 0.5 h

2.7 Read the luminescence intensity in ViewLux plate reader

3. Assay Performance

| PGC/ERR-agonist (Genistein) | Online Validation (Mean ± SD) |
| --- | --- |
| EC50 (µM) | 9.16±2.31 |
| S/B | 3.48±0.14 |
| CV (%) | 7.22±1.45 |
| Z’ | 0.65±0.08 |
| N | 27 |

| PGC/ERR-antagonist (XCT790) | Online Validation (Mean ± SD) |
| --- | --- |
| IC50 (µM) | 3.97±0.29 |
| S/B | 3.01±0.12 |
| CV (%) | 7.22± 1.45 |
| Z’ | 0.55± 0.07 |
| N | 27 |

| PGC/ERR-Viability (Tetra-Br.) | Online Validation (Mean ± SD) |
| --- | --- |
| IC50 (µM) | N/A |
| S/B | 7.64±0.23 |
| CV (%) | 3.96±0.85 |
| Z’ | 0.90±0.03 |
| N | 27 |